## Slide 1: Introduction to Quantifying STEM-EDXS (XEDS) data
Tom Slater

## Slide 2: Introduction
How do you go from a set of X-ray peaks to the composition of your sample?

## Slide 3: Introduction
X-ray generation
Sample specifics
Detector properties
Equation for X-ray intensity from sample in STEM-EDX

## Slide 4: X-ray generation
ionisation cross-section for α-shell of element A
number density of atoms of element A
incident number of electrons
sample thickness

## Slide 5: X-ray generation
Feldman & Mayer, Fundamentals of Surface and Thin Film Analysis
radiative partition function for α X-ray line
(proportion of transitions from the
particular sub-shell)
fluorescence yield (proportion of transitions
that result in X-ray emission (vs electron
emission))
for K X-rays
Kα1 = 100
Kα2 = 50
Kβ1 = 15-30
Kβ2 = 1-10
Kβ3 = 6-15

## Slide 6: Detecting X-rays
Current detector of choice – Silicon Drift Detector
Signal from FET

## Slide 7: Detecting X-rays
Detector efficiency for X-rays of
energy equivalent to the α X-ray
Collection solid angle

## Slide 8: Why are X-ray detectors always above sample? - Bremsstrahlung
Characteristic
~Isotropic
Continuum
Highly Anisotropic
Bremsstrahlung produced due to inelastic scattering of electron by nucleus (predominantly).
Bremsstrahlung Process
Bremsstrahlung is highly forward peaked.

## Slide 9: Bremsstrahlung continued
Kramers’ Law

## Slide 10: Quantification – Cliff-Lorimer
Cliff & Lorimer J. Microsc. (1975) 103, 203-207
Watanabe & Williams J. Microsc. (2006) 221, 89-109
Weight fraction of element A
Avogadro’s number
Density of sample
Atomic weight of element A
Cliff & Lorimer’s well known equation!
Don’t worry about this one!

## Slide 11: How do we get Cliff Lorimer k-factors?
Calculate
Measure
Need a reliable cross-section and detector efficiency in order to get an accurate answer.
Usually quoted to be accurate to ±10% for K peaks and ±20% for L peaks. Difficult for M peaks!
Get a standard sample and measure!
Your standard needs to:
Contain the elements you would like to measure.
Be a single known phase.
Not change composition due to electron beam damage.
Be sufficiently thin not to worry about absorption correction.
Can be tricky to get a suitable standard.

## Slide 12: Workflow for quantifying EDX/S data
Get spectrum (can be individual spectrum or many spectra as in a spectrum image).
Get intensity in each X-ray peak of interest with background subtracted.
Apply quantification using Cliff-Lorimer approach with known k-factors.

## Slide 13: Hyperspy – Getting Intensities

## Slide 14: Hyperspy – Getting Intensities

## Slide 15: Hyperspy - Quantification-K_Factors
Requires list of K-factors in same order as image stack.
K-factors based on Mass units.
Returns list of atomic (or weight) percent maps for each element.

## Slide 16: Errors in quantification
Error in Cliff-Lorimer factor
We have already mentioned that calculate k-factors may have an error of ±20%. If you want very accurate results you should always check with a standard but you will still be limited by X-ray statistics.
Error in X-ray intensity
You are measuring a Gaussian peak and therefore the standard deviation will be and the relative error with a 95% confidence level would be . Even if you have 10,000 counts in each peak that still gives a relative error of 2%.
Always think about the error in your quantification…

## Slide 17: Absorption Correction
t
STEM
Probe
Detector
z

z
For a thin film
)
Absorption along path l
Integrated absorption of all paths along the beam direction

## Slide 18: Quantification – Single element methods
Watanabe et al. Ultramic. (1996) 65, 187-198
Watanabe & Williams J. Microsc. (2006) 221, 89-109
K. E. MacArthur et al. Microsc. Microanal. 22 (2016) 71–81.
Zeta-factor
Cross-section

## Slide 19: Quantification-zeta_Factors
Requires list of zeta-factors in same order as image stack.
zeta-factors based on Mass units.
Returns a list of atomic (or weight) percent maps for each element and the mass thickness map for the sample.

## Slide 20: Quantification-cross_section
Requires list of experimental cross_sections in same order as image stack.
Cross sections based on atomic units.
Returns a list of atomic percent maps for each element.
Returns list of the number of atoms per pixel maps for each element.

## Slide 21: Quantification - Tips
Both zeta_factors and cross_sections require a calculation of the electron fluence. Therefore the following parameters must be set:
  s.set_microscope_parameters(beam_current = 0.0432)
  s.set_microscope_parameters(live_time = 0.003024)
If an element is repeated in the stack of images e.g. Ka and Kb it must be repeated in the list of factors. – This means you can specify a different quantification for each sub-shell if wanted.

## Slide 22: Limitations – (aka Topics we haven’t covered)
Have assumed thin film throughout, i.e. no fluorescence and no change in integrated flux with depth.
STEM
Probe
Primary X-ray
Secondary X-ray
X-ray fluorescence
In SEM have much more complicated matrix corrections.

## Slide 23: Limitations – (aka Topics we haven’t covered)
For X-ray absorption, assume a point detector but X-ray detectors keep subtending larger solid angles.
Should we integrate over a range of take-off angles?

## Slide 24: Limitations – (aka Topics we haven’t covered)
Collection solid angle may be fixed but ‘effective’ solid angle can change with holder / tilting of sample.
Important for spectra collected on E02 or if tilting on E01!